## Slide 1: mHealth and wearable devices
Theory and practice
Martin Úbl
2025

## Slide 2: Table of contents
Introduction, context, diabetes treatment
DIY vs. certified approach
mHealth and device requirements
Wearable devices and data collection
SmartCGMS

## Slide 3: Introduction and context
Diabetes
  Type 1, 2 and others
We focus on Type 1 Diabetes
  Autoimmune
  Manifests in early childhood
  Insulin treatment
  Currently requires a substantial amount of wearable electronics
https://diabetes.zcu.cz/

## Slide 4: Diabetes treatment
Insulin (Type 1, newly even for type 2)
  Insulin pen
  Insulin pump
    Subcutaneous
    Intradermal
Antidiabetic drugs (Type 2)

## Slide 5: Insulin dosing
Bolus
  Manual
Basal
  Manual
  Automatic
    How?

## Slide 6: Measurement
Glucose concentration
  In blood
    Glucometer
    Sporadic
  Subcutaneously
    CGM sensor
    „continuous“

## Slide 7: Typical „setup“
Sensor
Insulin pump
Infusion set
Controller (device)
  Optional
  Resides between the pump and the sensor
  Mobile phone, smart watch
Source: https://diabeteson.com/technical-devices-that-improve-risk-factors-care-and-quality-of-life/
Discussion: Lots of wearable devices, how does it influence patient‘s mental state (e.g., in children patients)?

## Slide 8: Closed control loop  artificial pancreas

## Slide 9: Controller development
Certified
  „the correct and safe one“
DIY
  „the immediatelly deployed one“

## Slide 10: DIY
Last 15-20 years
Patients themselves develop a treatment loop
  „Gluing together“ a number of components
  Algorithm prototyping
Is not a subject of certification
Risks vs. advantages?

## Slide 11: DIY in diabetes treatment
OpenAPS
  open-source, JavaScript, Python
  oref0 algorithm
AndroidAPS
  open-source, Java
  Runs the oref0 algorithm (JavaScript)
Loop
  iOS variant, Objective-C (later Swift)

## Slide 12: OpenAPS
mHealth?
Wearable electronics?
Safety?
  Old, deprecated devices
  JavaScript
  What if it fails?

## Slide 13: OpenAPS - failure
How do we detect and/or solve a failure?
OpenAPS code does not look like a safe code…
Console logging
No attempt of recovery after failure
…80 lines of code…

## Slide 14: OpenAPS – will not fail?
„OpenAPS cannot fail“
  Really?
  Statement supported by „tens of thousands of run-time“
We can partially avoid failure by verification
  OpenAPS has not been verified
  To be verifiable, the code needs to be prepared for it
    „spaghetti“ code of appx. 1600 lines of JavaScript certainly does not look like it is prepared

## Slide 15: AndroidAPS
Similar situation
  Java aplication for Android
  Runs JavaScript for insulin dose calculation
Author trusts his own software to a degree, when he set it up in a closed loop mode for his own daughter (10 years old)
The code is in a similar state, as the OpenAPS one
Author himself proclaims, that he is „not a good programmer“

## Slide 16: All systems
Wearable electronics
  Requires a communication protocol (network)
DIY systems use old, deprecated hardware
  A number of exploits in the protocol
  Buggy
  No warranty
  Who is responsible for injuries?
    Patient sets it up on his/her own
    A physician (diabetologist) tolerates the use, sometimes even encourages it

## Slide 17: DIY systems fail
In fact, pretty often
  A selection of recent issues in the OpenAPS repository
Contradictory control rules
Too much insulin due to sensor failure
Insufficient testing (CI/CD)
Bugs in code due to its unmaintainability
Silent failure of OpenAPS
Weird runtime requirements

## Slide 18: DIY systems fail
Hardware wears out, no continuous diagnostics (self-tests, POST, …)
Authors haven‘t considered timezones (local times) at all!
Used technologies are hard to maintain and prone to dependency-related errors

## Slide 19: According to what is DIY „safe“?
https://doi.org/10.1177/1932296818795705
https://doi.org/10.1111/dme.13816
https://openaps.org/2016/06/11/real-world-use-of-open-source-artificial-pancreas-systems-poster-presented-at-american-diabetes-association-scientific-sessions/
https://doi.org/10.2337/db18-993-P
https://doi.org/10.1177/1932296815583334
https://dx.doi.org/10.1177%2F1932296816665635

## Slide 20: Physicians love DIY
Most physicians only see the results
  Results are mostly good
Psychological aspect?
„ends justify the means?“, knowingly ignoring technical imperfections

## Slide 21: Requirements for mHealth devices
Algorithms are formally correct
  Verification
  Thorough testing within precisely bulit scenarios
    in-silico (pre-clinical)
    in-vivo (clinical)
Fault-tolerance and recovery
  Fault-tolerant properties
  Verification
Security
Lifecycle
  warranty, updates, regular technical maintenance, …

## Slide 22: Certification
FDA (USA), EMA (Europe)
Very difficult process
  Long
    Years of work (paperwork and additional work towards formal requirements)
  Expensive
    Even tens of millions $
  Laborious
Certified devices can guarantee certain degree of safety and correctness

## Slide 23: Classification of medical equipment (FDA)
Class I
  Minimal to no risk, do not directly affect patient‘s health
  E.g., fitness bands, thermometers, … even bandages and similar
Class II
  Moderate risks, may affect patient‘s health
  E.g., blood pressure meter, insulin pump (open-loop), glucometer, … even scalpels and needles
Class III
  High risk, affects patient‘s health, may cause serious injury or even death
  E.g., automatic insulin pump (closed-loop), CGM sensor, pacemaker, … even cochlear implant and joint replacement

## Slide 24: Equipment approval
Class I
  Register your product by the FDA
  Some exceptions may include additional paperwork
Class II
  Performance and effectivity is evaluated; even on market, they require collecting feedbacks and monitoring (for adverse effects and similar)
  Devices must have a unique serial number, patients must be registered
Class III
  Must undergo exhaustive testing and verification process
  Clinical studies with large number of participants
  Intentionally is a long process
    If there is a bug in device code or hardware, the longer time, the greater probability of failure

## Slide 25: Verification of algorithms and devices
Systematic testing of all possible states and validating responses
Simple example:
  Two-parametric controller
  Cartesian product of stepped parameter values in some (safe) boundaries
  Metric evaluation on a number of scenarios
  Attempt to identify „faulty“ combinations
  Alternatively, the implementation might be verified formally/mathematically (like the seL4 kernel)
  This is rarely possible

## Slide 26: Simple example
LGS – treatments standard, certified
BetaPID – adaptive PID controller
OpenAPS/oref1 – DIY
3 controllers, all having 2 parameters
Only 2 of them have visible safe regions
Legend:
  Red – probably lethal
  Yellow – edge case, potentially dangerous
  Green – the best the controller can do

## Slide 27: Wearable devices
Mobile phone, smart watch, fitness bands, but also CGM sensorts and more
A lot of sensors
  Lots of data
  Lots of possibilities
Accelerometer
Magnetometer
Ambient light sensor
GPS
Heartbeat sensor
Electrodermal activity sensor
Blood pressure sensor
Oxymeter
…

## Slide 28: Wearable devices - data
Personalized medicine?
  Treatment model personalization
Telemedicine?
  Physicians always have recent data
  A parent always sees recent data of his/her child
Development of new physiological/treatment models?
  Datasets for initial cross-validation

## Slide 29: Smart clothes
Not exactly a recent trend
Development of electronics-enhanced clothes
  Health monitoring
  Work assistance, safety
  Cool effects

## Slide 30: Wearable devices - data
Problems?

## Slide 31: SmartCGMS
Framework designed and implemented on our department
Signal analysis framework and architecture
Built in such a way, that it may reach production qualities
  Fault-tolerance
  Verifiability
  Simplicity
  Stability
  Multi-platform
  Effectivity, low-power
Supports simulations and real-time use

## Slide 32: SmartCGMS
Implementation split into modules of various types
  Filter
  Model
  Signal
  Solver
  Metric
  …
Every module can be verified separately
  Simplifies the verification process
New module = verification of a single module
  It is not necessary to verify the whole system
From simulation to real-world in just a few steps
  Matter of a single module swap

## Slide 33: SmartCGMS
Linear connection of filters
Message passing („from left to right“)
struct TDevice_Event {
NDevice_Event_Code event_code;
GUID device_id;
GUID signal_id;
double device_time;
int64_t logical_time;
uint64_t segment_id;
union {
double level;
IModel_Parameter_Vector* parameters;
wstr_container* information;
};
};

## Slide 34: SmartCGMS
Fulfills a role of the back-end – framework, set of components and SDK
Front-ends
  gpredict3 – science and development
  SmartCGMS Mobile – patient monitoring
  Games - Icarus has Diabetes, DiaAdventure
  Pump-Trainer – education of newly diagnosed patients
  Web simulator – WASM-based online simulator

## Slide 35: SmartCGMS – gpredict3

## Slide 36: SmartCGMS – Mobile

## Slide 37: SmartCGMS – Icarus has Diabetes

## Slide 38: SmartCGMS – DiaAdventure

## Slide 39: SmartCGMS – Pump-Trainer (old)

## Slide 40: SmartCGMS – Pump-Trainer (new)

## Slide 41: SmartCGMS – Web simulator

## Slide 42: Thank you for your attention
Questions, discussion, …
Martin Úbl
ublm@kiv.zcu.cz